# Integrating the Healthcare Enterprise

# IHE Radiation Oncology

Technical Framework Supplement

# Quality Assurance with Plan Veto Profile

# (QAPV)

# Draft in Preparation for Public Comment

Date:		Jan 26, 2016

Author:	Chris Pauer

Email:		chrispauer@sunnuclear.com

Foreword

This is a supplement to the IHE Radiation Oncology Technical Framework <VX.X>. Each supplement undergoes a process of public comment and trial implementation before being incorporated into the volumes of the Technical Frameworks.

<For Public Comment:> This supplement is published on <Month XX, 201x> for Public Comment. Comments are invited and may be submitted at http://www.ihe.net/<domain>/<domain>comments.cfm. In order to be considered in development of the Trial Implementation version of the supplement, comments must be received by <Month XX, 201X>.

This supplement describes changes to the existing technical framework documents.

“Boxed” instructions like the sample below indicate to the Volume Editor how to integrate the relevant section(s) into the relevant Technical Framework volume.

Amend section X.X by the following:

Where the amendment adds text, make the added text bold underline. Where the amendment removes text, make the removed text bold strikethrough. When entire new sections are added, introduce with editor’s instructions to “add new text” or similar, which for readability are not bolded or underlined.

General information about IHE can be found at: www.ihe.net.

Information about the IHE Radiation Oncology domain can be found at: http://www.ihe.net/Domains/index.cfm.

Information about the organization of IHE Technical Frameworks and Supplements and the process used to create them can be found at: http://www.ihe.net/About/process.cfm and http://www.ihe.net/profiles/index.cfm.

The current version of the IHE Radiation Oncology Technical Framework can be found at: http://www.ihe.net/Technical_Framework/index.cfm.

# CONTENTS

Introduction to this Supplement	7

Open Issues and Questions	7

Closed Issues	7

History	13

General Introduction	15

Appendix A - Actor Summary Definitions	15

Appendix B - Transaction Summary Definitions	15

Glossary	16

Volume 1 – Profiles	17

X Quality Assurance with Plan Veto Workflow Profile	18

X.1 QAPV Actors, Transactions, and Content Modules	18

X.1.1 Actor Descriptions and Actor Profile Requirements	20

X.1.1.1	Object Request Failure	20

X.1.1.2	Critical Issue	20

X.1.1.3	Plan Veto	20

X.2 QAPV Actor Options	21

X.3 QAPV Required Actor Groupings	21

X.4 QAPV Overview	21

X.4.1 Concepts	21

X.4.2 Use Cases	22

X.4.2.1 Use Case #1: Detect Dangerous Plan Specifications or Modifications	22

X.4.2.1.1 Detection Use Case Description	22

X.4.2.1.2 QAPV Process Flow	22

X.5 QAPV Security Considerations	25

X.6 QAPV Cross Profile Considerations	25

Appendices	26

Appendix A – Actor Summary Definitions	26

Appendix B – Transaction Summary Definitions	26

Volume 2 – Transactions	28

Y.1 RO-Q1: Create UPS for Quality Check	28

3.Y.1.1 Scope	28

Y.1.2 Actor Roles	28

Y.1.3 Referenced Standards	29

Y.1.4 Interaction Diagram	29

Y.1.4.1 N-Create Message	29

Y.1.4.1.1 Trigger Events	29

Y.1.4.1.2 Message Semantics	29

Y.1.4.1.2.1	Preconditions	29

Y.1.4.1.2.2	SOP InstanceUID	30

Y.1.4.1.2.3	Scheduled Procedure Step Start	30

Y.1.4.1.2.4	Scheduled Procedure Step Priority	30

Y.1.4.1.2.5	Input Readiness State	30

Y.1.4.1.2.6	Input Information Sequence Specification	30

Y.1.4.1.3 Expected Actions	31

Y.2	RO-Q2: Subscribe to UPS Progress Update	31

Y.2.1 Scope	31

Y.2.2 Actor Roles	31

Y.2.3 Referenced Standards	32

Y.2.4 Interaction Diagram	32

Y.2.4.1 Subscribe N-ACTION Message	32

Y.2.4.1.1 Trigger Events	32

Y.2.4.1.2 Message Semantics	32

Y.2.4.1.3 Expected Actions	33

Y.3 RO-Q3: Workitem Input Objects Retrieval  for Difference Check	33

Y.3.1 Scope	33

Y.3.2 Actor Roles	33

Y.3.3 Referenced Standards	34

Y.3.4 Interaction Diagram	34

Y.3.4.1 Workitem Objects Retrieval Message	34

Y.3.4.1.1	Preconditions	34

Y.3.4.1.2 Trigger Events	35

Y.3.4.1.3 Message Semantics	35

Y.3.4.1.3.1	Object Requirements for Evaluation [is this really enforcable…use of “shall”…]	35

Y.3.4.1.4 Expected Actions	38

Y.4 RO-Q4:Workitem Input Objects Retrieval for Dose Check	39

Y.4.1 Scope	39

Y.4.2 Actor Roles	39

Y.4.3 Referenced Standards	39

Y.4.4 Interaction Diagram	40

Y.4.4.1 Retrieve Workitem Objects Message	40

Y.4.4.1.1 Preconditions	40

Y.4.4.1.2 Trigger Events	40

Y.4.4.1.3 Message Semantics	41

Y.4.4.1.4 Expected Actions	42

Y.5 RO-Q5:Update on UPS Progress	42

Y.5.1 Scope	42

Y.5.2 Actor Roles	42

Y.5.3 Referenced Standards	43

Y.5.4 Interaction Diagram	43

Y.5.4.1 Update on UPS Progress	43

Y.5.4.1.1 Trigger Events	43

Y.5.4.1.1.1	Required Actions	43

Y.5.4.1.2 Message Semantics	44

Y.5.4.1.3 Expected Actions	44

Y.6 RO-Q6: Output Information Sequence Retrieval	44

Y.6.1 Scope	44

Y.6.2 Actor Roles	44

Y.6.3 Referenced Standards	45

Y.6.4 Interaction Diagram	45

Y.6.4.1 Output Information Sequence Retrieval	45

Y.6.4.1.1 Preconditions	45

Y.6.4.1.2 Trigger Events	45

Y.6.4.1.3 Message Semantics	46

Y.6.4.1.4 Expected Actions	47

Y.7 RO-Q7: Quality Check Results Retrieval	47

Y.7.1 Scope	47

Y.7.2 Actor Roles	47

Y.7.3 Referenced Standards	48

Y.7.4 Interaction Diagram	48

Y.7.4.1 Quality Check Results Retrieval Message	48

Y.7.4.1.1 Preconditions	48

Y.7.4.1.2 Trigger Events	49

Y.7.4.1.3 Message Semantics	49

Y.7.4.1.3 Expected Actions	50

Y.8 RO-Q8:Unsubscribe to UPS Progress	50

Y.8.1 Scope	50

Y.8.2 Actor Roles	50

Y.8.3 Referenced Standards	51

Y.8.4 Interaction Diagram	51

Y.8.4.1 Unsubscribe to UPS Progress Message	51

Y.8.4.1.1 Trigger Events	51

Y.8.4.1.2 Message Semantics	51

Appendices	52

Volume 2 Namespace Additions	52

Volume 3 – Content Modules	53

5. Namespaces and Vocabularies	54

6. Content Modules	55

7. DICOM Content Definition	56

7.3 IOD Definitions	56

7.3.2.2.4 RT Plan IOD for Content Assessment – Consistency Check	56

7.3.2.2.4.1 Referenced Standards	56

7.3.2.2.4.2 IOD Definition	56

7.3.2.2.5 RT Plan IOD for Content Assessment – Dose Check	57

7.3.2.2.5.1 Referenced Standards	57

7.3.2.2.5.2 IOD Definition	58

7.3.7.1.1 Content Assessment Results	 IOD for Quality Assurance with Plan Veto	60

7.3.7.1.1.1 Referenced Standards	60

7.3.7.1.1.2 IOD Definition	60

7.4 DICOM Content Modules	60

7.4.3.1.2 RT General Plan Module for Content Assessment – Consistency Check	61

7.4.3.2.4 RT Prescription Module for Content Assessment – Dose Check	62

7.4.3.3.4 RT Fraction Scheme Module for Content Assessment – Dose  Check	62

7.4.4.2.3 Control Point Attributes for Content Assessment – Dose Check	64

7.4.12 Content Assessment Module	65

Volume 4 – National Extensions	67

4 National Extensions	67

# Introduction to this Supplement

This supplement adds the Quality Assurance with Plan Veto (QAPV) profile to the IHE-RO Domain.  The QAPV Profile describes the interaction between a Quality Check Requester and a Quality Check Performer that will force evaluation of radiation treatment data to detect and avoid severe treatment errors.

## Open Issues and Questions

- Transaction numbers are suggested at this time and should be given appropriate numbers and letter codes.

- 60. Profile not correct in terms of DICOM supplement naming and content [Review]

61.	Handling of plan content: check if TD TP   BRTO can be used as reference to those existing modules. [Review]

62 Appropriate place to invoke adherence to TPPC / TDPC

63 Move Content Assessment content to a Content Module definition

64 Review requirement on QAPV_EQUIVALENT, shall be made precise and explicit…Is this appropriate here?

## Closed Issues

- [Closed]Until it is clear that the two types of safety checks can be handled under one profile, the current aim is to create a generic profile with the ability to extend it to more specific applications if the transactions and objects needed cannot be handled under one profile.  If multiple profiles are needed, section 2.1 below may need to be filled in.  Modified: 4/16/2011 – added Appendices to discuss specific formatting and data for different evaluations.  Will now try to cover both in this document, although the positioning evaluation will remain TBD for now.

Added Appendices to discuss specific formatting and data for different evaluations.  Will now try to cover both in this document, although the positioning evaluation will remain TBD for now

- [Closed] The name of this profile should be reviewed and commented on.

The name has been presented in various arena with no negative comments

- [Closed]Can we use the verification information objects defined in Supplement 74 to express the evaluation of information at risk in objects?  Do we have to limit this to plans?  Do they need to be extended in some way?

See issue 5 below.

- [Closed] How do the changes outlined in Beam Dose Depth DICOM CP affect this profile?  Additionally, the Beam Dose additions are associated at the dose reference sequence, not at the control point sequence, and so are not in line with what was intended.  Some implementations already have these in place.  Discussions continue with Christof, Bruce, Chris P, and Craig L.  [Update] Examples discussed in December to try to clarify how this can be specified for the profile.

- [Closed]Supplement 74 has a lot to say on this model, although there have been issues with the 74 approach brought up at WG-6.  These should be reviewed and taken into account.

Not investing in the machine verification model of Supplement 74, but much of the document is useful in framing the data interactions here.  Instead, the reporting of  what failed a test is currently following the recommendation of DICOM WG-7, of using a structured report.

- [Closed]We need more discussion on how the quality check rules are formulated, tested, reviewed and enforced.  A Quality Check Performer actor should comply with which of the following?  Section 1 in the supplement below will need to be modified depending on what we arrive at.

  - Level 1 of Rule Visibility: The quality check rules are considered to be fully under the control of application implementation, and so it is up to each vendor to develop, market and publicize how effective they are in implementing their safety checks.  Testing for profile compliance is limited to forcing a vetoable set of data to be sent, and a fail result being returned.

  - Level 2 of Rule Visibility: The actor needs to make the rules set viewable in a common way.  Testing and modification of the rules are closed to outside parties.  Contents of the rules are part of the application implementation and are the reponsibility of the vendor to develop and market how effective they are in implementing their safety checks. Compliance checking includes those tests in Level 1, and that the product also allows rule review in a common way.

  - Level 3 of Rule Visibility: Rules are viewable in a common way, and testing of the rules is structured and repeatable.  Rules can be added in defined, automated way to make the Quality checks more restrictive than the initial set that the vendor supplies.  Content of the initial rules is under complete control of the application vendor.  Testing for compliance would follow those under level 1 and 2, but also include tests to add and exercise new rules to the QA checks.

  - Level 4 of Rule Visibility: Rules are viewable in a common way, and testing of the rules is structured and repeatable.  The initial set of rules is defined  in a standard set as the responsibility of xxxx.  Actor compliance at this level means they will implement the rules that are defined in this standard, and structured testing will demonstrate that they force compliance to these checks or rules when operating as the Quality Check Performer.   Testing for compliance will use the standard set of rules to formulate data that can exercise the actor, as well as the tests referred to in level 1, 2 and 3.

Quality checks will be viewable and testable (as per the QA Advisory Group Position Statement)…but are going to be set by the clinical staff.  How the rules will be expressed or set is up to vendor control, but testing of critical check will be exercisable using RT Plan data.

- [Closed] Any other updates in RO-Q4 that need to be required?  None noted.

- [Closed] Should trigger events for the initial push of the Quality Check UPS be futher detailed?  None noted.

- [Closed]What value do checks other than those done ASAP have?  Should it be specified that other interested parties may want a check done, but not necessarily immediately before treatment?  Out of band for the profile.  Technically, the QCP will probably not be aware whether a check request is ASAP or not, unless it supports some scheduling options that are outside the profile purveyance.

- [Closed]Should the N-EVENT-REPORT updates be documented in further detail than they are in the high level sequence diagram for clarity?  N-Event Report will only be used to note the UPS has been completed/cancelled, not as a mechanism for reporting the success/failure of the check.

- [Closed]From Stuart:

In Table A.1 Required Input Sequence Content for Dosimetric QA 

I think that it should be made explicit that the contents of the objects need to adhere to the requirements made inother profiles. 

There is no current mechanism (profile) to enforce that the Treatment Delivery System will receive the list of input objects that are identified (other than RT Plan). The current profile is Treatment Delivery Workflow (TDW). It doesn't address imaging/positioning. 
When IPDW is available, that will, so CT and RT Structure Set will have a specific means of being expressed (perhaps that's true for TDW, but it is never required in TDW). However, not all Treatment Delivery Systems will necessarily support import of volumetric data (2D/2D imaging only), and for sure, not all imaging sessions that are part of a treatment session ("session" is being used in a loose sense here) will demand (be scheduled) volumetric data (CT and RT Structure Set).

One can make the argument that the RT Plan created by the TPS will have an RT Structure Set (required for the Basic RT Objects Interoperability Profile, 
and at least by inference in Advanced RT Objects Interoperability), so that even if it isn't an explicit part of the scheduled Treatment Delivery, the information is available to a device that needs to construct the Input Sequence Specification. And one can go even further and insist that because the RT Structure Set is referenced, a device can retrieve the object, inspect it, and identify which CT are required. 

But I don't see how one can ensure that the RT Dose will be uniquely referenced within the scope of an existing profile, nor how that RT Dose instance will be specified in the chain (from the Treatment Management System to the Treatment Delivery System). 
If it is necessary for the TMS to identify which RT Dose is involved "out of band", then that should be made explicit. 
(As a TMS vendor, I can see utility in having the TMS maintain the associations between RT Dose and RT Plan in a database).

However, there still isn't an explicit mechanism for getting this information over to the TDS utilizing a profile.

I suppose one can say that "it can go in the input sequence in TDW or IPDW", but that puts a burden on a device that is *not* in the profile. I don't know if RT Dose is actually needed for the use case (the QA vendors would need to weigh in). Beam Dose and Beam Meterset are part of the RT Plan.   If that *can* be enough, I believe that *should* be all that is required.

Chris’s note: The current set of attributes specified have been reviewed by those vendors expressing interest in creating a Quality Check Performer for Dose Check, and have been found to be sufficient.

- [Closed]MLC leaf opening data: This would seem to be important in calculating dose delivered to the patient, but for the purposes of this profile, two things should be considered: 1) leaf pattern is not a attribute of the RT Plan.  (At least the current generation) 2) Is it possible that the condition of a leaf that will be closed during a treatment would make the difference between a life-threatening delivery and a clearly non-life threatening delivery?  If the answer to 2 is “no”, then it seems we can ignore MLC leaf  data.  Are there other attributes missing from the DICOM objects in play that can make the calculation of critical levels of dose delivery be wildly inaccurate?  Leaf pattern IS part of the RT Plan.  Leaf Position Boundaries (300A,00BE)

- [Closed]Radiosurgery and hypofractionated treatment – Arguments against usefulness of this safety profile approach:

  - Hypofractionated doses might look like dangerous treatments and be flagged.

  - To exempt hypfractionated dose levels, the dose levels that are judged dangrous will be so high as to be:

    - Ludicrous

    - Never detected in any plan anywhere.

  - The RT industry doesn’t have the facility to reliably have a third party machine understand a structure whose dose limits are much lower (such as “spinal column” or “optic nerve”), and so extremely dangerous treatments can still slip through, and giving a false sense of security is really more dangerous than being clear that assessment and assurance needs to happen through a human agency every time.

The intent of the dose check component of this profile is to check for consistency between the dose described for coordinates, volume, site, or point, and the plan that has been formulated to deliver that dose.  It will not independantly judge dose level outside of the plan.  Under this scheme, a dose can be approved even if it is dangerously high if it is that way consistently in the plan, and matches what the plan delivers.   In other words, this check will not protect against bad prescriptions, it intends to check that the plan delivers what it was told to deliverIssue

- [Closed]The N-EVENT-REPORT usage refers to the Event Information parameter, however, I have difficulty finding a clear definition of how this is encoded. 
The best description I could find was in the Print SOP Class, where a set of DICOM elements were in fact the Event Information. 
If that is in fact the way that Event Information is encoded, then some choice needs to be made for the element or set of elements to use that constitute the Event Information.  

As a strawman, I can imagine using either the Approval Status module or the Substance Approval Module.   The N-Event report will not be used to express the check success/failure.

- [Closed]Insert a note regarding the device types that could fulfill the QCR actor role.

- [Closed] Check on addition of requirement for High Dose Technique (300A,00C7) in 3.4.5.5.  Remove, make equal to DICOM standard.  Atrribute removed from those listed as having requirements higher than DICOM standard.

- [Closed] Need new workitem codes defined for the two types of check requests.  RT reviewed and approved to go to WG-6 for CP assignment.

- [Closed] Review N-Create responses to make sure that appropriate responses are found for 3.1.5.3.7 and 3.1.5.3.8.   The response type that can be used for an unsupported or invalid Workitem Code Sequence Code Value would be Annex C.5.23 of Part 7, Unrecognized Operation.

- [Closed] Progression of states when cancelled. 3.3.5.1 and 3.4.5.1.  Progression state modified.

- [Closed] Review for any other required attributes of RT Plan for 3.4 and 3.6 Sent for approval to move forward with current attribute set 6/22/12

- [Closed] Decision on whether to include ION transactions or not.  Removed Ion Transactions from this track of the document (“A” editions)

- [Closed] Add precondition regarding readiness of report if stored in place other than the QCP itself

- [Closed] Object Store  Page 43, get rid of references to Object Store

- [Closed] Check that changes in Quality Check Report are reflected in rest of profile.

- [Closed] Need CP for templates in QAPV profile – RT 32

- [Closed] Need CP to add Quality Check Performer descriptor for Item 2 in template XXXX (main body of Check Result Report). – RT 33

- [Closed] Add “Critical” to item 6 in XXXX

- [Closed] Value Type should be “Text” for Item 6 (Num assumes a measurement and has units attached.)

- [Closed] Add “Critical” to item 7 in XXXX.

- [Closed] Duplicate number 9, make number 10

- [Closed] Edit item numbers of content items

- [Closed] Item 2 in TID YYYY (main body of structured report representing the Check Request report) should be code descriptor = IHE-RO.

- [Closed] Discuss items 1 and 2…they don’t seem to both need to be there or at all, if item 3 denotes the Issue Severity of this detail item entry

- [Closed] Item 3 should just be Issue is Critical  - Yes/No

- [Closed] Item 4 in YYYY, change value type to TEXT (more description, assume this is manufacturer code meaning…

- [Closed] Need CP for values in value set constraint for item 5 in YYYY – RT 34

- [Closed] Change item 5 in YYYY type to “CODE”

- [Closed] Various grammar and spelling issues

- [Closed] Add Output Info Seq return arrow on X.1-3

- [Closed] Question on end of 2.3.3 – TMS vs. TPS compliance.  Language allowing for TPS or TMS acting as a QCR is still in place in case a treatment delivery device is wholly controlled by one of these devices, and does not have the network or DICOM capability of managing the plan object itself.

- [Closed] Is Check Report available if the UPS is canceled?  “The Quality Check Performer is not required to create an instance of the Quality Check Report or to populate the Output Information Sequence of the Unified Procedure Step Performed Procedure Sequence if the step was CANCELLED”

- [Closed] R* in Object Requirements

- [Closed]  Reincorporate CP 1138

- [Closed]  ARTI requirements / RT Plan Requirements (Table 3.4) – No

- [Closed] Comment on Beam Limiting Device Sequence.  Any plan items that are actually part of the treatment should be included.  We will not force a beam limiting sequence if it is not neede for the treatment.

- [Closed] N-Get Tags and Requirements

- [Closed] Changes pending from Use Case review

- [Closed ] Statement on day / plan relationship and whether QCP is aware of this or checks it

- [Closed] Review TDW safety list for criteria to EXCLUDE a plan for assessment (patient name is not found, MRNID is not found…etc.)

- [Closed] Group review of updates for version 1.7A, based on discussion at november face-to-face meeting

- [Closed] Changes to definitions, transaction 3.3 as per December and January teleconferences.  Also, review outstanding changes to 3.7

- [Closed] Review changes on 1/17/13, to tie down difference check matching.  Also, change to structured report to allow other attributes for reporting critical issues

- [Closed] Changes for Mar 7, 2013 on Object Requirements for Evaluation

- [Closed] Object definition handled under DICOM supplement

- Remove material that will be in the DICOM standard.

- Added History tracking as used by DICOM WG-7 and other supplement authors

- Added display requirements for Quality Check Requester

- Added Notes on the Object Evaluation Results attributes.

## History

| Aug. 18, 2014 | 1.18 | Chris Pauer (cpauer@accuray.com) | Added History Table. |
| --- | --- | --- | --- |
| Aug. 19, 2014 | 1.18 | John Pauer (jcpauer@gmail.com) | Review and revise profile to shift references from Structured Report to Quality Assurance Results Object. |
| Aug. 21, 2014 | 1.18 | Chris Pauer (cpauer@accuray.com) | Review document against DICOM 2014 publication in preparation for presentation to QA Vendors review. |
| Oct 15, 2014 | 1.19 | Chris Pauer (cpauer@accuray.com) | Incorporate information from QA Dose call on Oct 14 |
| March 25, 2015 | 1.20 | Chris Pauer (cpauer@accuray.com) | Modified a few items to bring in line with Consistent Dose definition |
| April 20, 2015 | 1.21 | Chris Pauer (cpauer@accuray.com) | Further clean up, update references to the DICOM object used for reporting check results. |
| May 12, 2015 | 1.22 | Chris Pauer | Added attribute requirements for the Objet Evaluation Results |
| Oct 22, 2015 | 1.23 | Chris Pauer | Reviewed and changed for agreement with Public Comment version of Sup 185, and formatting as workflow profile. |
| Jan 26, 2016 | 1.24 | Chris Pauer | Edit for content modules, QA Plan handling, Cross profile concerns |

# General Introduction

Update the following Appendices to the General Introduction as indicated below. Note that these are not appendices to Volume 1.

# Appendix A - Actor Summary Definitions

Add the following actors to the IHE Technical Frameworks General Introduction list of Actors:

| Actor | Definition |
| --- | --- |
| Quality Check Requester | A system that will be performing treatment, or is in control of treatment on another device. In this role, it requests evaluation of treatment parameters. The evaluation, expressed as a pass or fail, is then accessible to this actor. This actor must implement Plan Veto behavior when the evaluation results in a failure. |
| Quality Check Performer | A device that evaluates treatment parameters. It compares the treatment attributes to an earlier quality checked version of the same plan, and/ or performs Critical Checks based on the comparison of Critical Values and Supplied Attribute Values. |

Appendix B - Transaction Summary Definitions

Add the following transactions to the IHE Technical Frameworks General Introduction list of Transactions:

| Transaction | Definition |
| --- | --- |
| RO-Q1: Create UPS for Quality Check | The Quality Check Performer implements the role of an SCP for UPS Worklist. The Quality Check Requester sends a request to create a UPS for a verification step. |
| RO-Q2: Subscribe to UPS Progress Update | The Quality Check Requester informs the Quality Check Performer that it wants updates on progress as it does the quality check. |
| RO-Q3: Workitem Input Objects Retrieval for Difference Check | The Quality Check Performer, using information supplied in the Input Information Sequence of the UPS from the Quality Check Requester, locates and retrieves data needed to do the difference evaluation. |
| RO-Q4: Workitem Input Objects Retrieval for Dose Check | The Quality Check Performer, using information supplied in the Input Information Sequence of the UPS from the Quality Check Requester, locates and retrieves data needed to do the dose evaluation. |
| RO-Q5: Update on UPS Progress | The Check Performer reports to the subscriber any important updates. |
| RO-Q6: Output Information Sequence Retrieval | Once the Quality Check Performer has signaled that the UPS is completed, the Quality Check Requester gets the Output Information Sequence of the UPS for the check. This location will describe where to find the Content Assessment Results Object containing the detailed information on the findings of the Quality Check and also contains the SOP Instance UID of the Results object. |
| RO-Q7: Quality Check Report Retrieval | After completion or cancellation of the UPS, the Quality Check Requester will perform a C-Move request for the Content Assessment Results detailing important information about the UPS processing, including the Critical Checks that were performed. The target of the Retrieval is found in the Output Information Sequence that was retrieved in transaction RO-Q6. (The Quality Check Performer is not required to supply a Content Assessment Results Object if the UPS status is CANCELLED, but it may be available) |
| RO-Q8: Unsubscribe to UPS Progress | The Quality Check Requester unsubscribes to the UPS updates from the Quality Check Performer. |

Glossary

Add the following glossary terms to the IHE Technical Frameworks General Introduction Glossary:

If these definitions are ambiguous when compared with existing items in the Glossary, please prepend “QAPV Profile” to the term.

| Glossary Term | Definition |
| --- | --- |
| Supplied Attribute | A specific property, quality or characteristic of data, expressed unambiguously in meaning and units, supplied from some known source. Also can refer to a sequence, set or construct consisting of individual properties, qualities or characteristics. |
| Calculated Attribute | A specific property, quality or characteristic of data, expressed unambiguously in meaning and units, that is calculate by an acknowledged source from a set of supplied or calculated attributes. Also can refer to a sequence, set or construct consisting of individual properties, qualities or characteristics. |
| Attribute Value | A particular entry, magnitude, number or amount of an Attribute. |
| Candidate Treatment Plan | A plan containing a description of device operation to be immediately used to treat a patient with radiation. This is the plan object retrieved from the Quality Check Requester to be used for the Quality Check. |
| Critical Value | A specific entry, magnitude, number or amount that is a notable or important marker in the allowable range of a specific Attribute Value. |
| Critical Check | A comparison between a Critical Value and an Attribute Value that will contribute to the safety assessment performed by a Quality Check Performer. |
| QA Assessed Plan | A treatment plan whose quality has been assessed and recorded by a QA Device that performs a Difference Check, and also checked by a clinical staff member. This assessment and recording happens prior to the transactions in the QAPV profile. The plan and check results are stored on the QA Device for later comparison with candidate treatment plans, and for constructing the resulting Results object. |
| QA Matched Plan | A single QA Assessed Plan that, through various criteria, has been found to be equivalent to the Candidate Treatment Plan. This plan is used to create the Content Assessment Results Object for the current invocation of the profile. |

Volume 1 – Profiles

Add the following to the IHE Technical Frameworks General Introduction Copyright section:

# Add to Section …

# X Quality Assurance with Plan Veto Workflow Profile

The Quality Assurance with Plan Veto Profile describes behavior and interactions to allow checking for critical safety issues prior to treatment.

## X.1 QAPV Actors, Transactions, and Content Modules

This section defines the actors, transactions, and/or content modules in this profile. General definitions of actors are given in the Technical Frameworks General Introduction Appendix A at http://www.ihe.net/Technical_Framework/index.cfm.

Figure X.1-1 shows the actors directly involved in the QAPV Profile and the relevant transactions between them. If needed for context, other actors that may be indirectly involved due to their participation in other related profiles are shown in dotted lines. Actors which have a mandatory grouping are shown in conjoined boxes.

# Figure X.1-1: QAPV Actor Diagram

Table X.1-1 lists the transactions for each actor directly involved in the QAPV Profile. To claim compliance with this Profile, an actor shall support all required transactions (labeled “R”) and may support the optional transactions (labeled “O”).

# Table X.1-1: QAPV Profile - Actors and Transactions

| Actors | Transactions | Optionality | Reference |
| --- | --- | --- | --- |
| Quality Check Performer | RO-Q1 Create UPS for Quality Check | R |  |
|  | RO-Q2 Subscribe to UPS Progress Update | R |  |
|  | RO-Q3 Workitem Input Objects Retrieval for Difference Check | RC – Required if RO-Q4 is not supported |  |
|  | RO-Q4 Workitem Input Objects Retrieval for Dose Check | RC – Required if RO-Q3 is not supported |  |
|  | RO-Q5 Update on UPS Progress | R |  |
|  | RO-Q6 Output Information Sequence Retrieval | R |  |
|  | RO-Q7 Quality Check Results Retrieval | R |  |
|  | RO-Q8 Unsubscribe to UPS Progress | R |  |
| Quality Check Requester | RO-Q1 Create UPS for Quality Check | R |  |
|  | RO-Q2 Subscribe to UPS Progress Update | R |  |
|  | RO-Q3 Workitem Input Objects Retrieval for Difference Check | RC – Required if RO-Q4 is not supported |  |
|  | RO-Q4 Workitem Input Objects Retrieval for Dose Check | RC – Required if RO-Q3 is not supported |  |
|  | RO-Q5 Update on UPS Progress | R |  |
|  | RO-Q6 Output Information Sequence Retrieval | R |  |
|  | RO-Q7 Quality Check Results Retrieval | R |  |
|  | RO-Q8 Unsubscribe to UPS Progress | R |  |

### X.1.1 Actor Descriptions and Actor Profile Requirements

Most requirements are documented in Transactions (Volume 2) and Content Modules (Volume 3). This section documents any additional requirements on profile’s actors.

#### X.1.1.1	Object Request Failure

If the object request made against the Quality Check Requester fails, either because the object is not found, or some other data inconsistency, the flow of interactions will follow those outlined in Sections 3.3.5.1 and 3.4.5.1: the state of the UPS will move from “SCHEDULED” to “IN PROGRESS” to “CANCELED” with appropriate updates being sent to the Quality Check Requester.  A CANCELED status means that the UPS could not be performed.  interactions after this point are outside the scope of this profile, but it does mean the check could not be performed, so the appropriate mitigation or exception handling is expected on the part of the Quality Check Requester actor.

#### X.1.1.2	Critical Issue

### X.1.1.3	Plan Veto

A critical requirement of this profile is that the Quality Check Requester relies upon the Quality Check Performer to deliver a pass/fail judgement on the safety of the plan.    The results of Critical Checks are communicated in the Content Assessment Results Object prepared by the Quality Check Performer.  If a critical issue is found, attribute Assessment Summary (GGGG, EE01) [pending inclusion in DICOM standard] in the Content Assessment Results will have the value “FAILED”.  The Quality Check Requester must not allow automatic treatment or automatic override if the value of this item is “FAILED”.  Only the Assessment Summary (GGGG,EE01) attribute value should be used to trigger Plan Veto.  Other entries in the Content Assessment Result are for informational use only.

      - The Quality Check must be performed at a point where treatment can be prevented.

      - The benefit of the Critical Check increases the closer in time that the Check occurs to the start of treatment.

Since the plan veto is a capability of the Quality Check Requester itself, and not subject to an interoperable transaction, the implementation of that capability is up to the actor vendor.  It is also expected that the Quality Check Requester will display in a clear and understandable way any critical results if the Content Assessment Results notes a Assessment Summary (GGGG, EE01) = FAILED.  The testing of the adherence to the profile will exercise and evaluate these capabilities.

The override, if implemented, is expected to be interactive in some engaging way, as opposed to a simple Yes/No button dialog.  It is suggested that the override present relevant portions of the Content Assessment Results and force entry and persistence of a senior clinical staff member’s id and password.  The information must be captured in an auditable fashion and this capability will be part of the profile testing.

Note also that even though it is strongly advised that the check happen as closely to, but previous to, treatment as possible, it may not be the Treatment Delivery Device that adheres to the Quality Check Requester actor.  The items required for the check may not be present on the Treatment Delivery Device, and so it may be required that the Treatment Management System, or Treatment Planning System has to fulfill the Quality Check Requester role.

#### X.2 QAPV Actor Options

#### Table X.2-1: Actor Options <<recast as cross profile dependencies>>

| Actor | Options | Details |
| --- | --- | --- |
| Quality Check Requester | Supports Treatment Delivery Plan Content | Supports production of content according to the Treatment Delivery Plan Content Definitions found in Volume 3 of the IHE-RO Technical Framework, as detailed in the product’s DICOM Conformance Statement. |
| Quality Check Performer | Supports Treatment Planning Plan Content | Supports consumption of content according to the Treatment Planning Plan Content Definitions found in Volume 3 of the IHE-RO Technical Framework, as detailed in the product’s DICOM Conformance Statement. <<for non supported plan types, should handle safely….see TMS support in TDPC / TPPC>> |
|  | Supports Treatment Delivery Plan Content | Supports consumption of content according to the Treatment Delivery Plan Content Definitions found in Volume 3 of the IHE-RO Technical Framework, as detailed in the product’s DICOM Conformance Statement. |

## X.3 QAPV Required Actor Groupings

There are no actor groupings for the QAPV profile.

## X.4 QAPV Overview

### X.4.1 Concepts

### X.4.2 Use Cases

#### X.4.2.1 Use Case #1: Detect Dangerous Plan Specifications or Modifications

##### X.4.2.1.1 Detection Use Case Description

This profile grew out of a request from ASTRO to address patient safety via an IHE-RO profile.  As such, it does not address a current clinical failing or use case that is commonly found, but tries to set up an inter-device check of plan attributes that can detect treatment issues.  It cannot do it narrowly, and cannot fix a bad prescription.

The profile also implements a transaction flow that can be used to automate general treatment machine to QA device checks.

##### X.4.2.1.2 QAPV Process Flow

# Figure X.4.2.2-1: Basic Process Flow in QAPV Profile

Pre-conditions:

The profile assumes some Critical Values have been defined on the Quality Check Performer.  If the Critical Values have not been defined, the Quality Check Performer should accept the request to create a Unified Procedure Step to do a Quality Check, and will be required to cancel it after an initial subscription has been received.  The profile also allows for the Quality Check Performer to indicate other required steps before the interactions described here can be supported.

Because of the varied Critical Checks and requirements of different Quality Check Performer products, it would be prudent for Quality Check Requesters to allow for configuration of, and submission to multiple Quality Check Performers for each treatment.  It is anticipated that Quality Check Performers will implement models of performing checks which may not be appropriate for all patients at a given site.

Main Flow:

The Quality Assurance with Plan Veto profile describes behavior and interactions to augment safety checks during a patient’s treatment in radiation therapy:

- The QAPV Check Performer accepts a UPS Workitem for a verification action from a Check Requester

- The Check Requester subscribes to updates on the workitem.

- The Check Performer actor requests the data specified in the UPS Workitem as required to perform the evaluation.  If the data is not available, or deemed insufficient, the evaluation is not performed.

- The Check Performer performs the check and prepares the resulting results object for retrieval by the Check Requester.

- The Check Performer updates subscribers on the work that was done, including final state.

- The Check Requester gets the Output Information Sequence from the UPS maintained by the Check Performer.

- The Check Requester requests retrieval of a results object from the Check Performer using the SOP Instance and AE Title specified in the Output Information Sequence.

- The Check Requester unsubscribes to the workitem.

Post-conditions:

The Check Requester allows or prevents the treatment to proceed based on the information found in the Content Assessment Results Object.  Specific requirements on the treatment approval and override behavior can be found in section X.1.1.2

## X.5 QAPV Security Considerations

## X.6 QAPV Cross Profile Considerations

| Quality Assurance with Plan Veto | Treatment Delivery Plan Content (TDPC) | Candidate Treatment Plan or QA Assessed Plan (Beam) types defined in TDPC that are supported by a Quality Check Performer shall be documented in the Quality Check Performer’s Conformance Statement <<direct on how to list which content is supported>> IHE-RO Technical Framework content adherence can be noted in the Real World Application notes.[?} [Does not fall under Specialized Sop Class, nor a standard extended SOP class] | It is anticipated that different beam models will have different requirements and safety checks, and using the TDPC to drive the attribute requirements is appropriate to insure this occurs correctly. |
| --- | --- | --- | --- |
| Quality Assurance with Plan Veto | Treatment Planning Plan Content (TPPC) | QA Assessed Plan (Beam) types defined in TPPC that are supported by a Quality Check Performer should be documented in their Conformance Statement | It is anticipated that different beam models will have different requirements and safety checks, and using the TPPC to drive the attribute requirements is appropriate to insure this occurs correctly. |

Appendices

# Appendix A – Actor Summary Definitions

Quality Check Requester – A system that will be performing treatment, or is in control of treatment on another device.  In this role, it requests evaluation of treatment parameters.  The evaluation, expressed as a pass or fail, is then accessible to this actor.  This actor must implement Plan Veto behavior when the evaluation results in a failure.

Quality Check Performer – A device that evaluates treatment parameters.  It performs Critical Checks based on the comparison of Critical Values and Supplied Attribute Values.

# Appendix B – Transaction Summary Definitions

# RO-Q1: Create UPS for Quality Check

The Quality Check Performer implements the role of an SCP for UPS Worklist.  The Quality Check Requester sends a request to create a UPS for a verification step.

# RO-Q2: Subscribe to UPS Progress Update

In this transaction, the Quality Check Requester informs the Quality Check Performer that it wants updates on progress as it does the quality check.

# RO-Q3: Workitem Input Objects Retrieval for Difference Check

In the Workitem Input Objects Retrieval for Difference Check transaction, the Quality Check Performer, using information supplied in the Input Information Sequence of the UPS from the Quality Check Requester, locates and retrieves data needed to do the difference evaluation.

# RO-Q4: Workitem Input Objects Retrieval for Dose Check

In the Workitem Input Objects Retrieval for Dose Check transaction, the Quality Check Performer, using information supplied in the Input Information Sequence of the UPS from the Quality Check Requester, locates and retrieves data needed to do the dose evaluation.

# RO-Q5: Update on UPS Progress

In the Update on UPS Progress transaction, the Check Performer reports to the subscriber any important updates.

# RO-Q6: Output Information Sequence Retrieval

Once the Quality Check Performer has signaled that the UPS is Completed, the Quality Check Requester gets the Output Information Sequence of the UPS for the check.   This location will describe where to find the results object containing the detailed information on the findings of the Quality Check and also contains the SOP Instance UID of the results object.

# RO-Q7: Quality Check Report Retrieval

After Completion or Cancellation of the UPS, the Quality Check Requester will perform a C-Move request for the results object detailing important information about the UPS processing, including the Critical Checks that were performed.  The target of the Retrieval is found in the Output Information Sequence that was retrieved in transaction RO-Q6.  (The Quality Check Performer is not required to supply a Results Object if the UPS status is CANCELED, but it may be available)

# RO-Q8: Unsubscribe to UPS Progress

In this transaction, the Quality Check Requester unsubscribes to UPS updates from the Quality Check Performer.

Volume 2 – Transactions

# Add section Y

## Y.1 RO-Q1: Create UPS for Quality Check

### 3.Y.1.1 Scope

In the Create UPS for Quality Check transaction, a Quality Check Performer receives a request from a Quality Check Requester to create a UPS with an appropriate verification workitem type.

### Y.1.2 Actor Roles

# Figure Y.1.2-1: Use Case Diagram

# Table Y.1.2-1: Actor Roles

| Actor: | Quality Check Requester |
| --- | --- |
| Role: | Uses the N-CREATE DIMSE service with a UPS Push SOP class to create a procedure to perform a quality check on the Quality Check Performer |
| Actor: | Quality Check Performer |
| Role: | Receives request for the UPS Push of a UPS Worklist object from a Quality Check Requester and creates the Unified Procedure Step |

Transaction text specifies behavior for each Role. The behavior of specific Actors may also be specified when it goes beyond that of the general Role.

### Y.1.3 Referenced Standards

# DICOM Standard 2011  PS 3.4-2011 Annex CC.2.1

### Y.1.4 Interaction Diagram

#### Y.1.4.1 N-Create Message

This is a UPS Push worklist message sent to the Quality Check Performer.

##### Y.1.4.1.1 Trigger Events

Either automatically triggered by treatment workflow, or by a specific user gesture on the Quality Check Requester, or one of the devices it is controlling.

##### Y.1.4.1.2 Message Semantics

The Quality Check Requester uses the N-CREATE request of the DICOM Unified Procedure Step – Push SOP Class to push a workitem onto the worklist of the Quality Check Performer. The Quality Check Requester performs the SCU role, and the Quality Check Performer performs the SCP role.  The message contains a formatted UPS under the UPS Push SOP class indicating the desired evaluation.

###### Y.1.4.1.2.1	Preconditions

The Quality Check Performer can stipulate  requirements of its choosing before accepting requests for creating a Quality Check procedure step.   In general, these requirements are meant to cover configuration steps or other critical settings that need to be in place before a quality check can be performed.  These requirements are outside the scope of this profile.  The requirements stipulated cannot violate the specific details or general intention of the profile.

The Quality Check Performer should specify if it needs to limit active subscribed UPSs for Quality Check to some number at any given time.

##### Y.1.4.1.2.2	SOP InstanceUID

This value is required by the standard in the SOP Common Module Attributes, but the DICOM Standard states that in an N-CREATE request message, it is “not allowed”.    This should be left unassigned, and is the task of the Quality Check Performer to assign, and return a valid SOP Instance UID.

##### Y.1.4.1.2.3	Scheduled Procedure Step Start

The earliest Scheduled Procedure Step Start Date and Time (0040,4005) of the UPS can be used by the Quality Check Performer to set the order of checks if processing more than one Quality Check Request at a time.

##### Y.1.4.1.2.4	Scheduled Procedure Step Priority

Quality Check Performer should in general ignore this attribute, and process check requests in the order received.

##### Y.1.4.1.2.5	Input Readiness State

Shall be READY.  This indicates that the Workitem needed for the Quality Check is available for retrieval by the Performer via C-MOVE

##### Y.1.4.1.2.6	Input Information Sequence Specification

The Input Information Sequence (0040,4021) shall contain reference to one of the following items :

The RT Plan SOP Instance to be delivered. The specified location should contain the plan in a state exactly as it is planned to be treated.

# Table Y.1  Required Input Sequence Content for Quality Check

| SOP Class Name | SOP Class UID | Condition |
| --- | --- | --- |
| RT Plan Storage | 1.2.840.10008.5.1.4.1.1.481.5 | If transaction RO-Q3 or RO-Q4 is supported |

Y.1.4.1.2.7	Required Attribute Values

The Scheduled Workitem Code Sequence (0040,4018) Code Value shall be either value 121731 (RT Treatment QA with RT Plan Dose Check) or 121732 (RT Treatment QA with RT Plan Difference Check) under Context ID 9241 in the DICOM standard, and Coding Scheme Designator shall be equal to ‘DCM’.

# Y.1.4.1.2.8	Return Value

If the N-CREATE request cannot be handled by the Quality Check Performer, because of missing references, invalid message formatting, invalid workitem code or other reasons, an error response should be returned as per DICOM Part 7 Annex C.  Because the subscription request has not been received from the Check Requester at this point, there is no other facility to signal a canceled workitem at this point.  Canceling a request because of the existence of a prior UPS for Quality Check is not acceptable up to the limit of concurrent checks the Quality Check Performer has specified.

The successful response should contain the SOP Instance UID of the Unified Procedure Step created in this transaction.

##### Y.1.4.1.3 Expected Actions

# Quality Check Performer creates the expected Unified Procedure Step

## Y.2	RO-Q2: Subscribe to UPS Progress Update

### Y.2.1 Scope

In the Subscribe to UPS Progress Update transaction, the Quality Check Requester issues a request to the Quality Check Performer to receive updates on the evaluation.

### Y.2.2 Actor Roles

# Figure Y.2.2-1: Use Case Diagram

# Table Y.2.2-1: Actor Roles

| Actor: | Quality Check Requester |
| --- | --- |
| Role: | Uses the N-ACTION service request under a UPS Watch SOP class to subscribe to the UPS created in transaction RO-Q1 |
| Actor: | Quality Check Performer |
| Role: | Processes subscription request |

Transaction text specifies behavior for each Role. The behavior of specific Actors may also be specified when it goes beyond that of the general Role.

### Y.2.3 Referenced Standards

# DICOM Standard 2011  PS 3.4-2011 Annex CC.2.3

### Y.2.4 Interaction Diagram

#### Y.2.4.1 Subscribe N-ACTION Message

##### Y.2.4.1.1 Trigger Events

After successful notification from the Quality Check Performer that the UPS has been created, the Quality Check Requester should issue this transaction.

##### Y.2.4.1.2 Message Semantics

This message should specify the UPS Watch SOP Class as the abstract syntax, but specify the original UPS Push SOP Instance UID of the UPS created in RO-Q1 as the Requested SOP Instance UID in the N-ACTION message.

Y.2.4.1.2.1	Global and Locking Subscriptions

The Quality Check Performer should refuse Global subscription requests. Only non-locking subscriptions should be processed.

##### Y.2.4.1.3 Expected Actions

Quality Check Performer subscribes the requester to the updates of the Unified Procedure Step.

## Y.3 RO-Q3: Workitem Input Objects Retrieval  for Difference Check

### Y.3.1 Scope

In the RO-Q3 Workitem Input Objects Retrieval for Difference Check transaction, the Quality Check Performer requests and receives from the target specified in the UPS in RO-Q1  the RT Plan instance required for performing the evaluation.   For the Difference evaluation, this instance of the RT Plan is the Candidate Treatment Plan.

### Y.3.2 Actor Roles

# Figure Y.3.2-1: Use Case Diagram

# Table Y.3.2-1: Actor Roles

| Actor: | Quality Check Requester |
| --- | --- |
| Role: | Receives request and sends requested DICOM objects to the Quality Check Performer. |
| Actor: | Quality Check Performer |
| Role: | Requests and receives requested DICOM objects from the Quality Check Requester |

Transaction text specifies behavior for each Role. The behavior of specific Actors may also be specified when it goes beyond that of the general Role.

### Y.3.3 Referenced Standards

# DICOM Standard 2011  PS 3.4: Storage Service Class

DICOM Standard 2011 PS 3.4: Query/Retrieve Service Class

### Y.3.4 Interaction Diagram

#### Y.3.4.1 Workitem Objects Retrieval Message

##### Y.3.4.1.1	Preconditions

It is expected that for Quality Check Performer actors that implement the Difference Check transaction some previous assessment or approval has been applied to zero or more RT Plans before the transactions in this profile have been started.  As a result of these previous checks, a set of QA Assessed Plans will be stored on the Quality Check Performer.  The Candidate Treatment Plan moved in this transaction will be compared against a subset of the QA Assessed Plans to determine if a QA Matched Plan exists.

It is also expected that the version of the plan that is assessed by this QA device, and that which is used for treatment will differ.  If this is the case, the Quality Check Performer which implements the Difference Check, at the time of the pre-profile assessment,  should expect to receive a full treatment plan and also a modified treatment plan that it will actually do the assessment on.  It will need to store the results of the assessment along with the modified treatment plan and full treatment plan in order to fully comply with the behavior of this transaction.

The Retrieve (Study Root – MOVE) SOP Class shall be supported. Implementations shall support modes of operation in which a single series (e.g. input CT Series) or specific SOP Instances (e.g. an RT Plan) are retrieved from the Quality Check Requester using the Study Root – MOVE SOP Class. Refer to DICOM 2007 PS 3.4, Annex C, for detailed descriptive semantics.

A Quality Check Performer shall be capable of issuing Study-Root C-MOVE for the RT Plan Object.

A Quality Check Performer  may receive an RT Plan in the Input Information Sequence for which it determines that it cannot perform the Difference Check.  (Situations where this can occur are discussed below in “Y.3.4.1.3.1 Object Requirements for Evaluation”)

In such cases:

- If the Procedure Step is not yet “IN PROGRESS”, the Quality Check Performer should move the Procedure Step to “IN PROGRESS” and then to “CANCELED” and update any subscribers.

- If the Procedure Step is already “IN PROGRESS”, the Quality Check Performer shall move the Procedure step to “CANCELED” and update any subscribers

##### Y.3.4.1.2 Trigger Events

The Quality Check Performer, in order to perform a Quality Check, requests the RT Plan referenced in the Input Information Sequence (0040,4021) from the Quality Check Requester.

##### Y.3.4.1.3 Message Semantics

The message semantics are defined by the DICOM Query/Retrieve SOP Classes and the DICOM Object Storage SOP Classes.

A C-MOVE Request from the DICOM Study Root Query/Retrieve Information Model – MOVE SOP Class shall be sent from the Quality Check Performer to the Quality Check Requester.

#### Y.3.4.1.3.1	Object Requirements for Evaluation [is this really enforcable…use of “shall”…]

Refer to Volume 3 – 7.4.3.1.2 General Plan Module for Content Assessment – Consistency Check for requirements on the incoming Candidate Treatment Plan.

The incoming Candidate Treatment Plan should be compared to the plans that were previously assessed by the Quality Check Performer, called the QA Assessed Plans.  This comparison is expected to be done to verify that an earlier, out of profile scope, quality check is still valid.

The rules for matching a Candidate Treatment Plan have been formulated with the understanding that the following events are common clinical practice.  These practices may not comply with the exact requirements of the DICOM specification, but products in the space should expect to encounter them, and the structure of the profile accounts for this.

- Patient name or ID being changed/coerced in a plan without generating a new SOP instance UID for the plan (“Bob”->”Robert”)

- Non-treatment beams being added or removed from a plan after export from the TPS.

- Delivery parameters being changed outside of the TPS (which typically results in the plan no longer being dosimetrically equivalent)

Every time that a plan is changed and the corresponding new SOP Instance UID is created for the modified plan, one of two things shall happen:

- If the change does not affect any dosimetric parameters, then the new plan shall have a Referenced RT Plan Sequence added that points to the SOP Instance UID of the previous plan.  The RT Plan Relationship value shall be QAPV_EQUIVALENT to denote that these two plans shall be assumed equivalent for the purposes of the safety checks in this profile.

- If the change does affect dosimetric parameters, then the new plan shall have all Referenced RT Plan Sequences with RT Plan Relationship value set to QAPV_EQUIVALENT removed, as the two plans are no longer equivalent for the purposes of the profile.

When the Quality Check Performer receives a Candidate Treatment Plan, every QA Assessed Plan that can be linked to the Candidate Treatment Plan must be collected and compared for dosimetric equivalence.  A QA Assessed Plan can be linked to a Candidate Treatment Plan by the SOP Instances matching directly, the SOP Instance of one appearing in the QAPV_EQUIVALENT list of the other, or by having a SOP Instance common to both QAPV_EQUIVALENT lists.

If there are zero linked plans in the set of QA Assessed Plans, then the Quality Check Performer must cancel the request.

If any of the linked QA Assessed Plans are not dosimetrically equivalent to the Candidate Treatment Plan, that QA Assessed Plan becomes the QA Matched Plan and the structured report must indicate Major Issue Found, due to the possibility that the Candidate Treatment Plan has been changed in a way that is dangerous to the patient.  Otherwise, the plans are considered equivalent, and the QA Matched Plan is chosen by checking:

- If any of the linked plans have Major Issue Found set to YES, then the most recent of those is chosen to be QA Matched Plan, and the structured report must indicate a veto.

- Otherwise, the most recent linked plan is the QA Matched Plan, and the structured report is constructed accordingly.

# Figure Y.3.3 QAPV_EQUIVALENT Processing

For the purposes of this example, the plans with SOP Instance UID 3.3.3 ,4.4.4, and 5.5.5 have been previously assessed by the Quality Check Performer.  These plans are related via QAPV Equivalence as indicated in the image above.

The Candidate Treatment Plan is one of the two plans with SOP Instance UID 2.2.2, which has 1.1.1 in the QAPV Equivalent list, also as indicated above.

The Candidate Treatment Plan is linked to both 3.3.3 and 4.4.4 due to all three plans having 1.1.1 in their QAPV Equivalent lists.

If any dosimetric differences exist between 2.2.2 and 3.3.3, or between 2.2.2 and 4.4.4, then the Candidate Treatment Plan will be vetoed by returning a structured report with Major Issue Found set to YES.

The rules for how a product curates the set of QA Assessed Plans is out of scope of this profile.  It is expected that the Quality Check Performer will have a set of QA Assessed Plans with more than one member (for example, one or more plan per patient currently undergoing treatment), but that is not required by the profile.  It is also expected, but not required, that the set of QA Assessed Plans will include plans that have negative quality assessments, so that the Quality Check Performer can inform the Quality Check Requester that the plan should be vetoed, rather than returning a Cancel due to lack of information.

The Quality Check Performer needs to be able to clearly communicate to the users what plans are present in the set of QA Assessed Plans, as well as if a given plan check is “more recent” than another.  The method by which this information is communicated is out of scope for this profile.

For the purposes of plan comparison, how delivery parameters should be compared is not explicitly defined by the profile, and is left to the discretion of each implementation of the QCP actor. In general, it is expected that plans that do not have appreciable differences in their delivery parameters should have equivalent treatment effect upon delivery as each other.

The profile does not require other values to be checked, such as day and time of treatment.  Since the emphasis of the profile currently is to catch the most egregious cases, it is felt that this is not ideal, but passable.  The Candidate Treatment Plan will be compared against a previously checked and passed plan, whether it is the one appropriate to today’s treatment or not.  This does not preclude this check from being part of a specific implementation.

The class of plans that a Quality Check Performer will be able to evaluate shall be specified in the DICOM Conformance statement in terms of the  beam types and options within the ARTI profile.

The Quality Check Performer is expected to document any exceptions to the supported ARTI Beam Models that would result in a CANCELATION of the requested assessment.  In general, it should be possible to infer which plans a Quality Check Performer can assess from its DICOM Conformance Statement.

##### Y.3.4.1.4 Expected Actions

The Quality Check Requester receives the C-MOVE request, establishes a DICOM association with the Quality Check Performer actor, and transfers the requested object.

The Quality Check Performer is then expected to use the requested objects in the performing of the desired evaluation.

## Y.4 RO-Q4:Workitem Input Objects Retrieval for Dose Check

### Y.4.1 Scope

In the Workitem Input Objects Retrieval for Dose Check transaction, the Quality Check Performer requests and receives from the Quality Check Requester the RT Plan specified in the UPS in RO-Q1  required for performing the evaluation.  The Plan instance must have been supplied in the Input Information Sequence of the initial UPS N-CREATE request.

### Y.4.2 Actor Roles

# Figure Y.4.2-1: Use Case Diagram

# Table Y.4.2-1: Actor Roles

| Actor: | Quality Check Requester |
| --- | --- |
| Role: | Receives C-MOVE request and sends requested DICOM objects to the Quality Check Performer |
| Actor: | Quality Check Performer |
| Role: | Requests and receives requested DICOM objects from the Quality Check Requester. |

Transaction text specifies behavior for each Role. The behavior of specific Actors may also be specified when it goes beyond that of the general Role.

### Y.4.3 Referenced Standards

# DICOM Standard 2011  PS 3.4:Storage Service Class

DICOM Standard 2011 PS 3.4: Query/Retrieve Service Class

### Y.4.4 Interaction Diagram

#### Y.4.4.1 Retrieve Workitem Objects Message

This is a UPS Push worklist message sent to the Quality Check Performer.

##### Y.4.4.1.1 Preconditions

The Retrieve (Study Root – MOVE) SOP Class shall be supported. Implementations shall support modes of operation in which a single series (e.g. input CT Series) or specific SOP Instances (e.g. an RT Plan) are retrieved from the Quality Check Requester using the Study Root – MOVE SOP Class. Refer to DICOM 2011 PS 3.4, Annex C, for detailed descriptive semantics.

A Quality Check Performer shall be capable of issuing Study-Root C-MOVE for the RT Plan specified in the Input Information Sequence. Other mechanisms for obtaining the data (such as C-STORE or restoring from a DICOM medium) shall not be relied upon to obtain the data.

A Quality Check Performer  may receive SOP Instance UIDs in the Input Information Sequence for which it determines that it cannot perform the Quality Check safely. In such cases:

- If the Procedure Step is not yet “IN PROGRESS”, the Quality Check Performer should move the Procedure Step to “IN PROGRESS” then to “CANCELED” and update any subscribers.

- If the Procedure Step is already “IN PROGRESS”, the Quality Check Performer shall move the Procedure Step to “CANCELED” and update any subscribers

##### Y.4.4.1.2 Trigger Events

The Quality Check Performer, in order to perform a Quality Check, requests the RT Plan referenced in the Input Information Sequence (0040,4021) that it needs for the selected work item.

##### Y.4.4.1.3 Message Semantics

The message semantics are defined by the DICOM Query/Retrieve SOP Classes and the DICOM Object Storage SOP Classes.

A C-MOVE Request from the DICOM Study Root Query/Retrieve Information Model – MOVE SOP Class shall be sent from the Quality Check Performer to the location specified in the UPS Input Information Sequence.

It is assumed that the RT Plan has been readied for retrieval by a means outside the scope of this profile.

# Y.4.4.1.3.1	Object Requirements for Evaluation

The RT Plan that is moved in transaction RO-Q4 has requirements under this profile that are more restrictive than the DICOM standard.

Transaction Requirements that are the same as the DICOM requirements are not listed below, unless it helps in the readability of these tables.

References to Notes or other DICOM sections have been deleted.  The DICOM standard should be referenced for supporting information, and for other attribute requirements.

The class of plans that a Quality Check Performer will be able to evaluate shall be specified in the DICOM Conformance statement in terms of the  beam types and options within the ARTI profile.

The Quality Check Performer is expected to document any exceptions to the supported ARTI Beam Models that would result in a CANCELATION of the requested assessment.  In general, it should be possible to infer which plans a Quality Check Performer can assess from its DICOM Conformance Statement.

All extant DICOM specifications in the RT Plan should also be supplied to the Quality Check Perfomer.  The Quality Check Performer should get an RT Plan that is at least as fully detailed as what the Quality Check Requester requires to do treatment, and beyond that, there are further more stringent requirements for this case as documented below.

| RT Plan Module / Attributes | Technical Framework Location |
| --- | --- |
| RT Prescription Module | Volume 3 7.4.3.2.4 |
| RT Fraction Scheme Module | Volume 3 7.4.3.3.4 |
| Control Point Attributes | Volume 3 7.4.4.2.3 |

##### Y.4.4.1.4 Expected Actions

The Quality Check Requester receives the C-MOVE request, establishes a DICOM association with the Quality Check Performer actor, and uses the appropriate DICOM Object SOP Classes to transfer the requested objects.

The Quality Check Performer is expected to consume the plan and to judge, based on required values in the RT Plan object in comparison with the Critical Values that were previously defined, whether the RT Plan, when delivered, would result in an intolerable radiation exposure.   Specifically, the Quality Check Performer should take the RT Prescription specifications, and using the beam description supplied in the RT Beams Module, make a judgment of whether the beam delivery will exceed the prescription in a radical way.

## Y.5 RO-Q5:Update on UPS Progress

### Y.5.1 Scope

In the Update on UPS Progress transaction, the Quality Check Performer updates all subscribed actors on any important attribute changes to the UPS.

### Y.5.2 Actor Roles

# Figure Y.5.2-1: Use Case Diagram

# Table Y.5.2-1: Actor Roles

| Actor: | Quality Check Performer |
| --- | --- |
| Role: | Sends N-EVENT-REPORT to update the Quality Check Requester subscriber when a UPS change occurs. |
| Actor: | Quality Check Requester |
| Role: | Receives N-EVENT-REPORT on updates to subscribed UPS. |

Transaction text specifies behavior for each Role. The behavior of specific Actors may also be specified when it goes beyond that of the general Role.

### Y.5.3 Referenced Standards

# DICOM Standard 2011  PS 3.4-2011 Annex CC.2.4

### Y.5.4 Interaction Diagram

#### Y.5.4.1 Update on UPS Progress

The Quality Check Performer uses the N-EVENT-REPORT of the UPS Event SOP Class to inform the Quality Check Requester of important changes to the UPS.  Any state change is appropriate, but there is one required update, documented below.

##### Y.5.4.1.1 Trigger Events

When important changes occur to the UPS while the Quality Check Performer is running the evaluation, N-EVENT-REPORT messages should be sent.

#### Y.5.4.1.1.1	Required Actions

The Quality Check Performer MUST send an N-EVENT-REPORT when the UPS is completed or discontinued.

If the evaluation is judged to be undoable, the UPS should be set to CANCELED, and this N-EVENT-REPORT should be sent to the Check Requester.

##### Y.5.4.1.2 Message Semantics

In either a COMPLETED or CANCELED state change, status report should come with Event Type ID 1. (Table C.C.2.4-1 in PS 3.4-2011).  If the UPS is set to CANCELED, the Reason for Cancellation and Procedure Step Discontinuation Reason Code Sequence are required under this profile.  Defined Context ID for the Code Sequence is 9300.

##### Y.5.4.1.3 Expected Actions

When completion is signaled, the Quality Check Requester should perform the RO-Q6 transaction.

## Y.6 RO-Q6: Output Information Sequence Retrieval

### Y.6.1 Scope

In the Output Information Sequence Retrieval transaction, the Quality Check Requester issues a request to the Quality Check Performer to get the Output Information Sequence of the Quality Check UPS.

### Y.6.2 Actor Roles

# Figure Y.6.2-1: Use Case Diagram

# Table Y.6.2-1: Actor Roles

| Actor: | Quality Check Requester |
| --- | --- |
| Role: | Uses the N-GET service request under a UPS Watch SOP class to retrieve attributes of the UPS created in transaction RO-Q1 |
| Actor: | Quality Check Performer |
| Role: | Honors the request. |

Transaction text specifies behavior for each Role. The behavior of specific Actors may also be specified when it goes beyond that of the general Role.

### Y.6.3 Referenced Standards

# DICOM Standard 2011  PS 3.4-2011 Annex CC.2.7

### Y.6.4 Interaction Diagram

#### Y.6.4.1 Output Information Sequence Retrieval

This is a UPS Push worklist message sent to the Quality Check Performer.

##### Y.6.4.1.1 Preconditions

The location of the Quality Check Results is indicated by the information fetched in RO-Q6.  This location is not required to be the AETitle and Address of the Quality Check Performer itself. The Quality Check Performer must insure that the instance of the Content Assessment Result is stored and available for retrieval before issuing RO-Q6.

##### Y.6.4.1.2 Trigger Events

After successful notification from the Quality Check Performer that the UPS has been  Completed or Canceled, the Quality Check Requester should issue this transaction.  The Quality Check Performer is not required to populate the Output Information Sequence if the Procedure was CANCELED.

##### Y.6.4.1.3 Message Semantics

This message should specify the UPS Watch SOP Class as the abstract syntax, but specify the original UPS Push SOP Instance UID of the UPS created in RO-Q1 as the Requested SOP Instance UID in the N-GET message.

If the UPS returns Procedure Step State (0074,1000) with value of CANCELED, the Quality Check Requester can inspect the Progress Information Sequence for the Reason for Cancellation (0074,1238) and Procedure Step Discontinuation Reason Code Sequence (0074,100E) for information on the Cancelation.  Defined Context ID for the Code Sequence is 9300.  Actions taken after determination of the Cancelation reasons are outside the scope of this profile. The Quality Check Performer is not required to create an instance of the Content Assessment Result or to populate the Output Information Sequence of the Unified Procedure Step Performed Procedure Sequence if the step was CANCELED.  In the case of a CANCELED step, the Quality Check Performer can choose to populate an Content Assessment Result, and in that case, needs to populate the Output Information Sequence with the location of the report.

If the UPS returns Procedure Step State (0074,1000) with a value of COMPLETED, the Quality Check Requester will request the value  of the Unified Procedure Step Performed Procedure Sequence, and get the values contained in the Output Information Sequence (0040,4033).

The Quality Check Requester can request attributes as follows:

| Attribute Name | Tag | Type (SCU/SCP) | Additional Information |
| --- | --- | --- | --- |
| Unified Procedure Step Progress Information Module |
| Procedure Step State | (0074,1008) | R+*/1 |  |
| Progress Information Sequence | (0074,1000) | R+*/R+* | Will include Cancelation information on return, if UPS was CANCELED. |
| Unified Procedure Step Performed Procedure Information Module |
| Unified Procedure Step Performed Procedure Sequence | (0074,1216) | R+*/R+* | Shall be returned if exists. |

All other standard N-Get behavior should be supported, so the above table in no way limits the information that the Quality Check Requester inquire about.

##### Y.6.4.1.4 Expected Actions

The Quality Check Requester is expected to retrieve the Content Assessment Result object if it exists.

## Y.7 RO-Q7: Quality Check Results Retrieval

### Y.7.1 Scope

In the Quality Check Results Retrieval transaction, a Quality Check Requester retrieves a Content Assessment Results Object using the location obtained in RO-Q6.  The results object indicates the checks that were performed and the status of each.

### Y.7.2 Actor Roles

# Figure Y.7.2-1: Use Case Diagram

# Table Y.7.2-1: Actor Roles

| Actor: | Quality Check Performer |
| --- | --- |
| Role: | Accepts request to move a DICOM Content Assessment Results Object to the Quality Check Requester. |
| Actor: | Quality Check Requester |
| Role: | Requests and receives the requested Content Assessment Results Object |

Transaction text specifies behavior for each Role. The behavior of specific Actors may also be specified when it goes beyond that of the general Role.

### Y.7.3 Referenced Standards

# DICOM 2011 PS 3.4: Storage Service Class

DICOM 2011 PS 3.16

DICOM 2011 PS3.3 A.35

DICOM 2011 PS 3.4: Query/Retrieve Service Class

### Y.7.4 Interaction Diagram

#### Y.7.4.1 Quality Check Results Retrieval Message

This is a UPS Push worklist message sent to the Quality Check Performer.

##### Y.7.4.1.1 Preconditions

The Retrieve (Study Root – MOVE) SOP Class shall be supported by the Quality Check Performer as an SCP.  Implementations shall support modes of operation in which a single series (e.g. input CT Series) or specific SOP Instances (e.g. an RT Plan) are retrieved from the Quality Check Performer using the Study Root – MOVE SOP Class. Refer to DICOM 2007 PS 3.4, Annex C, for detailed descriptive semantics.  The Quality Check Performer shall support the Query / Retrieve Service on instance level.

A Quality Check Requester shall be capable of issuing a Study-Root C-MOVE to obtain Content Assessment Results Object created by the Quality Check Performer

The C-STORE service for Content Assessment Results ( SOP Class UID 1.2.840.10008.5.1.4.1.1.xx.y.z) will be supported by the Quality Check Requester as an SCP.

If Critical Issues were found, the Quality Check Performer must prepare a Content Assessment Results Object.

Volume 3, Section 7.4.12 lists the attributes for the Content Assessment Results Object that have more specific requirements than what is detailed in the DICOM Standard.

##### Y.7.4.1.2 Trigger Events

The Quality Check Requester has retrieved the Output Information Sequence from the Quality Check Performer in transaction RO-Q6

##### Y.7.4.1.3 Message Semantics

The message semantics are defined by the DICOM Query/Retrieve SOP Classes and the DICOM Object Storage SOP Classes.

A C-MOVE Request from the DICOM Study Root Query/Retrieve Information Model – MOVE SOP Class shall be sent from the Quality Check Requester to the Quality Check Performer location specified in the Output Information Sequence

It is assumed, when the UPS is reported as “COMPLETED” that the Content Assessment Results Object has been fully prepared before the Quality Check Performer does the final state update via transaction RO-Q5

A Quality Check Performer that implements the Dose Check should create an Content Assessment Results Object with Assessment Summary set to FAILED when the Quality Check Performer determines that delivering the Candidate Treatment Plan would pose a threat to the health or well-being of the patient.

A Quality Check Performer that implements the Difference Check should create an Content Assessment Results Object with Assessment Summary set to FAILED when the Candidate Treatment Plan matches one or more QA Assessed Plan and:

- A single matching QA Assessed Plan was not approved in the earlier quality check.

- There are multiple matching QA Assessed Plans, and at least one of them was failed in the earlier quality check.

- The Candidate Treatment Plan has changed in some way that may affect the dosemetric delivery or safe treatment of the patient.

The Quality Check Requester must behave according to the guidelines in section X.1.1.2 “Plan Veto” when a Assessment Summary is FAILED.

The SOP class expected to be moved:

| Quality Check Performer (SCU) SOP Class | SOP Class UID |
| --- | --- |
| Content Assessment Results | [1.2.840.10008.5.1.4.1.1.xx.y.z] |

##### Y.7.4.1.3 Expected Actions

If the results object indicates a critical issue was found, the Critical Check Requester must veto delivery in a clear manner.

## Y.8 RO-Q8:Unsubscribe to UPS Progress

### Y.8.1 Scope

In the Unsubscribe to UPS Progress transaction, the Quality Check Requester issues a request to the Quality Check Performer to stop receiving updates on the evaluation.

### Y.8.2 Actor Roles

# Figure Y.8.2-1: Use Case Diagram

# Table Y.8.2-1: Actor Roles

| Actor: | Quality Check Requester |
| --- | --- |
| Role: | Uses the N-ACTION service request under a UPS Watch SOP class to unsubscribe to the UPS created in transaction RO-Q1 |
| Actor: | Quality Check Performer |
| Role: | Honors the request. |

Transaction text specifies behavior for each Role. The behavior of specific Actors may also be specified when it goes beyond that of the general Role.

### Y.8.3 Referenced Standards

# DICOM Standard 2011  PS 3.4-2011 Annex CC.2.3

### Y.8.4 Interaction Diagram

#### Y.8.4.1 Unsubscribe to UPS Progress Message

##### Y.8.4.1.1 Trigger Events

After successful notification from the Quality Check Performer that the UPS has been canceled, or the evaluation is complete, the Quality Check Requester should issue this transaction.

##### Y.8.4.1.2 Message Semantics

This message should specify the UPS Watch SOP Class as the abstract syntax, but specify the original UPS Push SOP Class UID of the UPS created in RO-Q1 as the Requested SOP Instance UID in the N-ACTION message.

The Quality Check Requester should implement this behavior in all cases to be well behaved.

Appendices

# Volume 2 Namespace Additions

Add the following terms to the IHE General Introduction Appendix G:

<There will be namespace additions pending a DICOM change proposal to add codes and the structured report items.>

Volume 3 – Content Modules

# 5. Namespaces and Vocabularies

Add to section 5 Namespaces and Vocabularies

# 6. Content Modules

Add to sections 6.4 and 6.5 Value Sets

# 7. DICOM Content Definition

## 7.3 IOD Definitions

##### 7.3.2.2.4 RT Plan IOD for Content Assessment – Consistency Check

###### 7.3.2.2.4.1 Referenced Standards

DICOM PS 3.3

###### 7.3.2.2.4.2 IOD Definition

| IE | Module | Reference | Usage | IHE-RO Usage |
| --- | --- | --- | --- | --- |
| Patient | Patient | C.7.1.1 | M | M See 7.4.1.1.1 |
|  | Clinical Trial Subject | C.7.1.3 | U | U |
| Study | General Study | C.7.2.1 | M | M See 7.4.1.2.1 |
|  | Patient Study | C.7.2.2 | U | U |
|  | Clinical Trial Study | C.7.2.3 | U | U |
| Series | RT Series | C.8.8.1 | M | M See 7.4.1.4.1 |
|  | Clinical Trial Series | C.7.3.2 | U | U |
| Frame of Reference | Frame of Reference | C.7.4.1 | U | - |
| Equipment | General Equipment | C.7.5.1 | M | M See 7.4.1.5.1 |
| Plan | RT General Plan | C.8.8.9 | M | R See 7.4.3.1.2 RT General Plan Module for Content Assessment – Consistency Check. Must also comply with 7.4.3.1.1 |
|  | RT Prescription | C.8.8.10 | U | Consistent with Appropriate Plan Content for Planning or Delivery State |
|  | RT Tolerance Tables | C.8.8.11 | U | Consistent with Appropriate Plan Content for Planning or Delivery State |
|  | RT Patient Setup | C.8.8.12 | U | R See 7.4.5.3.1 |
|  | RT Fraction Scheme | C.8.8.13 | U | Consistent with Appropriate Plan Content for Planning or Delivery State |
|  | RT Beams | C.8.8.14 | C - Required if RT Fraction Scheme Module exists and Number of Beams (300A,0080) is greater than zero for one or more fraction groups | Consistent with Appropriate Plan Content for Planning or Delivery State |
|  | RT Brachy Application Setups | C.8.8.15 | C - Required if RT Fraction Scheme Module exists and Number of Brachy Application Setups (300A,00A0) is greater than zero for one or more fraction groups | Absent |
|  | Approval | C.8.8.16 | U | R |
|  | SOP Common | C.12.1 | M | M See 7.4.1.6.1 |

##### 7.3.2.2.5 RT Plan IOD for Content Assessment – Dose Check

###### 7.3.2.2.5.1 Referenced Standards

DICOM PS 3.3

###### 7.3.2.2.5.2 IOD Definition

| IE | Module | Reference | Usage | IHE-RO Usage |
| --- | --- | --- | --- | --- |
| Patient | Patient | C.7.1.1 | M | M See 7.4.1.1.1 |
|  | Clinical Trial Subject | C.7.1.3 | U | U |
| Study | General Study | C.7.2.1 | M | M See 7.4.1.2.1 |
|  | Patient Study | C.7.2.2 | U | U |
|  | Clinical Trial Study | C.7.2.3 | U | U |
| Series | RT Series | C.8.8.1 | M | M See 7.4.1.4.1 |
|  | Clinical Trial Series | C.7.3.2 | U | U |
| Frame of Reference | Frame of Reference | C.7.4.1 | U | - |
| Equipment | General Equipment | C.7.5.1 | M | M See 7.4.1.5.1 |
| Plan | RT General Plan | C.8.8.9 | M | Consistent with Appropriate Plan Content for Planning or Delivery State |
|  | RT Prescription | C.8.8.10 | U | R See 7.4.3.2.4 RT Prescription Module for Content Assessment – Dose Check |
|  | RT Tolerance Tables | C.8.8.11 | U | Consistent with Appropriate Plan Content for Planning or Delivery State |
|  | RT Patient Setup | C.8.8.12 | U | R See 7.4.5.3.1 |
|  | RT Fraction Scheme | C.8.8.13 | U | R See 7.4.3.3.4 RT Fraction Scheme Module for Content Assessment – Dose Check |
|  | RT Beams | C.8.8.14 | C - Required if RT Fraction Scheme Module exists and Number of Beams (300A,0080) is greater than zero for one or more fraction groups | R Use appropriate Beams Module to Plan Type – Include 7.4.4.2.3 Control Point Attributes for Content Assessment – Dose Check |
|  | RT Brachy Application Setups | C.8.8.15 | C - Required if RT Fraction Scheme Module exists and Number of Brachy Application Setups (300A,00A0) is greater than zero for one or more fraction groups | Absent |
|  | Approval | C.8.8.16 | U | R |
|  | SOP Common | C.12.1 | M | M See 7.4.1.6.1 |

##### 7.3.7.1.1 Content Assessment Results	 IOD for Quality Assurance with Plan Veto

###### 7.3.7.1.1.1 Referenced Standards

DICOM PS 3.3

###### 7.3.7.1.1.2 IOD Definition

| IE | Module | Reference | Usage | IHE-RO Usage |
| --- | --- | --- | --- | --- |
| Patient | Patient | C.7.1.1 | M | M |
|  | Clinical Trial Subject | C.7.1.3 | U | U |
| Study | General Study | C.7.2.1 | M | M |
|  | Patient Study | C.7.2.2 | U | U |
|  | Clinical Trial Study | C.7.2.3 | U | U |
| Series | General Series | C.7.3.1 | M | M |
|  | Clinical Trial Series | C.7.3.2 | U | U |
| Equipment | General Equipment | C.7.5.1 | M | M |
|  | Enhanced General Equipment | C.7.5.2 | M | M |
| Content Assessment Results | Content Assessment Results | C.8.8.x | M | M |
|  | SOP Common | C.12.1 | M | M |
|  | Common Instance Reference | C.12.2 | M | M |

## 7.4 DICOM Content Modules

<Authors’ instructions: DICOM content modules will alter the expected behavior of DICOM modules and attributes as specified in DICOM Part 03.  It is strongly recommended that the author(s) review the structure of DICOM Part 03 and generally follow its format in the information presented in this section.   A critical understanding of DICOM definitions for cardinality, optionality, coded terminology values, and DICOM module structure, as well as IHE formatting conventions is also necessary. It is strongly recommended that the author is also conversant with the IHE Technical Frameworks General Introduction Appendix E “Conventions”.>

<For each DICOM content entry, there will be a sub-section labeled with a content acronym when added to the Technical Framework.

##### 7.4.3.1.2 RT General Plan Module for Content Assessment – Consistency Check

##### [Should not conflict with 7.4.3.1.1 – Checked…does not conflict ]

# Table 7.4.3.1.2.X Required Attributes for RT General Module for Content Assessment – Consistency Check in RT Plan

| Attribute Name | Tag | DICOM Type | Transaction Req. | Attribute Description |
| --- | --- | --- | --- | --- |
| Referenced RT Plan Sequence | (300C,0002) | 3 | R+* | Related instances of RT Plan. One or more items are permitted in this sequence. |
| >Referenced SOP Class UID | (0008,1150) | 1 | R* | Uniquely identifies the referenced SOP Class. |
| >Referenced SOP Instance UID | (0008,1155) | 1 | R* | Uniquely identifies the referenced SOP Instance |
| RT Plan Relationship | (300A,0055) | 1 | R+* | Relationship of referenced plan with respect to current plan. Defined Terms: PRIOR = plan delivered prior to current treatment ALTERNATIVE = alternative plan prepared for current treatment PREDECESSOR = plan used in derivation of current plan QAPV_EQUIVALENT = referenced plan is purported to be dosemetrically equivalent to the current plan VERIFIED_PLAN = plan which is verified using the current plan. This value shall only be used if Plan Intent (300A,000A) is present and has a value of VERIFICATION CONCURRENT = plan that forms part of a set of two or more RT Plan instances representing a single conceptual ‘plan’, applied in parallel in one treatment phase |

##### [Below does not conflict with CDEB]7.4.3.2.4 RT Prescription Module for Content Assessment – Dose Check

# Table 7.4.3.2.4.X Required Attributes for RT Prescription Module for Content Assessment – Dose Check in RT Plan

| Attribute Name | Tag | DICOM Type | Transaction Req. | Attribute Description |
| --- | --- | --- | --- | --- |
| Dose Reference Sequence | (300A,0010) | 3 | R+* | Any dose references used in Referenced Dose Reference Sequence (300C, 0050) in the RT Beams Module must be included in this sequence. This sequence shall contain at least 1 entry of Dose Reference Type (300A,0020) = TARGET |

##### 7.4.3.3.4 RT Fraction Scheme Module for Content Assessment – Dose  Check

# Table 7.4.3.3.4.X Required Attributes for RT Fraction Scheme Module for Content Assessment – Dose Check in RT Plan

| Attribute Name | Tag | DICOM Type | Transaction Req. | Attribute Description |
| --- | --- | --- | --- | --- |
| Fraction Group Sequence | (300A,0070) | 1 | - | Introduces sequence of Fraction Groups in current Fraction Scheme. One or more items shall be included in this sequence. |
| >Fraction Group Number | (300A,0071) | 1 | - | Identification number of the Fraction Group. The value of Fraction Group Number (300A,0071) shall be unique within the RT Plan in which it is created. |
| >Referenced Dose Sequence | (300C,0080) | 3 | - | Related instances of RT Dose (for grids, isodose curves and named/unnamed point doses). One or more items are permitted in this sequence. See Note 1. |
|  | >>Include ‘SOP Instance Reference Macro' Table 10-11 |
| >Number of Fractions Planned | (300A,0078) | 2 | R+* [TPPC/CDEB have R+] | Total number of treatments (Fractions) prescribed for current Fraction Group. Required to allow safety assessment of treatment course. |
| >Referenced Beam Sequence | (300C,0004) | 1C | - | Introduces sequence of treatment beams in current Fraction Group. One or more items shall be included in this sequence. Required if Number of Beams (300A,0080) is greater than zero. |
| >>Beam Dose Specification Point | (300A,0082) | 3 | R+* | Coordinates (x,y,z) of point at which Beam Dose is specified in the patient based coordinate system described in C.7.6.2.1.1 (mm). |
| >>Beam Dose | (300A,0084) | 3 | R+* More restrictive requirements may be in place for attribute in Planning or Delivery State | Dose (in Gy) at Beam Dose Specification Point (300A,0082) due to current Beam. |
| >>Beam Dose Verification Control Point Sequence | (300A,008C) | 3 | R+* [TPPC has R+] | Sequence of Items containing Beam Dose Verification Control Points. |
| >>>Cumulative Meterset Weight | (300A,0134) | 1 | - | The cumulative Meterset weight value, at which the beam dose point geometrical parameters apply. |
| >>>Referenced Control Point Index | (300C,00F0) | 1C | - | Uniquely identifies the Control Point specified by Control Point Index (300A,0112) within Beam referenced by Referenced Beam Number (300C,0006). Required, if the Cumulative Meterset Weight corresponds to a Control Point in the Control Point Sequence (300A,0111). |
| >>>Average Beam Dose Point Depth | (300A,008D) | 2C | - | The average depth (in mm) in the patient along a ray from the source to the dose point specified by the Beam Dose Specification Point (300A,0082) from the current Beam Dose Verification Control Point to the next one. |
| >>>Average Beam Dose Equivalent Depth | (300A, 008E) | 2C | - | The average radiological depth in mm (water-equivalent depth, taking tissue heterogeneity into account) in the patient along a ray from the source to the dose point specified by the Beam Dose Specification Point (300A,0082) from the current Beam Dose Verification Control Point to the next one. Required for all but the last items in that sequence. |
| >>>Average Beam Dose Point SSD | (300A,008F) | 2C | - | Average Source to patient surface distance in mm along a ray from the source to the dose point specified by the Beam Dose Specification Point (300A,0082) from the current Beam Dose Verification Control Point to the next one. Required for all but the last items in that sequence. |
| >>Beam Meterset | (300A,0086) | 3 | R+* | Machine setting to be delivered for current Beam, specified in Monitor Units (MU) or minutes as defined by Primary Dosimeter Unit (300A,00B3) (in RT Beams Module) for referenced Beam. |

##### 7.4.4.2.3 Control Point Attributes for Content Assessment – Dose Check

# Table 7.4.4.2.3.X Control Point Attributes for Content Assessment – Dose Check in RT Plan

| Attribute Name | Tag | DICOM Type | Transaction Req. | Attribute Description |
| --- | --- | --- | --- | --- |
| Beam Sequence | (300A,00B0) | 1 | - | Introduces sequence of treatment beams for current RT Plan. One or more items shall be included in this sequence. |
| >Primary Fluence Mode Sequence | (3002,0050) | 3 | R+* | Sequence defining whether the primary fluence of the treatment beam uses a non-standard fluence-shaping. Only a single item is permitted in this sequence. Shall be present. |
| >Control Point Sequence | (300A,0111) | 1 | - | Introduces sequence of machine configurations describing treatment beam. Two or more items shall be included in this sequence. |
| >>Control Point Index | (300A,0112) | 1 | - | Index of current Control Point, starting at 0 for first Control Point. |
| >>Referenced Dose Reference Sequence | (300C,0050) | 3 | RC* | Introduces a sequence of Dose References for current Beam. One or more items are permitted in this sequence. This sequence should be populated for all beams that will have a dose delivery effect on the dose references of type TARGET listed in the RT Prescription Module |
| >>>Referenced Dose Reference Number | (300C,0051) | 1 | - | Uniquely identifies Dose Reference specified by Dose Reference Number (300A,0012) in Dose Reference Sequence (300A,0010) in RT Prescription Module. |
| >>>Cumulative Dose Reference Coefficient | (300A,010C) | 2 | R+* [Same as CDEB] | Coefficient used to calculate cumulative dose contribution from this Beam to the referenced Dose Reference at the current Control Point. Must be present and non-null |

##### 7.4.12 Content Assessment Module

##### 7.4.12.1   Content Assessment Results Module for Quality Assurance with Plan Veto

Table 7.4.12.1-X Content Assessment Results Module Attributes for Quality Assurance with Plan Veto

Note: Attributes not shown should follow the DICOM standard type and description.

| Attribute Name | Tag | DICOM Type | Profile Req. | Attribute Description |
| --- | --- | --- | --- | --- |
| Assessment Summary | (gggg,ee01) | 1 | R+ | If a finding indicates that there may be severe effects to the patient, or if the plan material differs in a significant way from the assessed plan, then this value shall be FAILED |
| Assessment Type Code Sequence | (gggg,ee21) | 1 | R+* | Type of Assessment that was performed. |
| >Code Value | (0008,0100) | 1C | R+* | For a Dose Check use XYZ001. For a Difference/Consistency Check, use XYZ002 |
| >Coding Scheme Designator | (0008,0103) | 1C | R+* | Shall be DCM |
| >Code Meaning | (0008,0104) | 1 | D | Value shall be RT Pre-Treatment Dose Check if Code Value (0008,0100) = XYZ001. Value shall be RT Pre-Treatment Consistency Check if Code Value (0008,0100)=XYZ002 |
| Assessment Requester Sequence | (gggg,ee17) | 2 | R+* | The person or device that made the assessment request. Only a single item shall be included in this Sequence. |
| >Include C.17.2.4 Identified Person or Device Macro |
| Assessed SOP Instance Sequence | (gggg,ee04) | 1 | R+* | References to the SOP Instances that were supplied for assessment. Only one item shall be included in this sequence. |
| >Include Table C.17-3 ‘Hierarchical SOP Instance Reference Macro Attributes’ | R+ | SOP Class UID shall indicate RT Plan Storage |
| >Referenced Comparison SOP Instance Sequence | (gggg,ee05) | 1C | RC+* | Referenced SOP Instance used by the assessor to compare to the assessed SOP Instances. Required if the assessor used persistent data to compare to the assessed SOP Instance. Only one item shall be included in this sequence |
| >Include Table C.17-3 ‘Hierarchical SOP Instance Reference Macro Attributes’ | RC+* | SOP Class UID shall indicate RT Plan Storage |
| Number of Assessment Observations | (gggg,ee06) | 1 | - |  |
| Assessment Observations Sequence | (gggg,ee07) | 1C | RC+ | Observations made during the assessment. Required if Number of Assessment Observations (gggg,ee06) is not zero. The number of Items included in the Sequence shall equal the value of Number of Assessment Observations (gggg,ee06). |
| >Observation Type | (gggg,ee08) | 1 | R+ | Display required if value = “MAJOR” |
| >Observation Description | (gggg,ee0A) | 1 | R+ | Display required if value of level of concern = “MAJOR”. This description of the issue should be complete enough to allow the user to determine whether the fault can be overridden. |

Volume 4 – National Extensions

# Add appropriate Country section

# 4 National Extensions

Not applicable.

# Text boxes

Workitem Input Objects Retrieval for Dose [RO-Q4]

Quality Check Requester

Quality Check Performer

Create UPS for Quality Check

[

RO-Q1]

Subscribe to UPS Progress Update

[RO-Q2]

Update on UPS Progress

[RO-Q5]

Work item Input Objects Retrieval for Difference [RO-Q3]

→

Unsubscribe to UPS Progress [RO-Q8]

←

→

←

Quality Check Results Retrieval [RO-Q7]

←

→

←

Output Information Sequence Retrieval [RO-Q6]   ←

Create UPS for Quality Check [RO-Q1]

Perform

Evaluation

Quality Check Performer

Quality Check Requester

Subscribe to UPS Progress Update [RO-Q2]

Update on UPS Progress [RO-Q5]

Repeat

Unsubscribe to UPS Progress [RO-Q8]

alt

Workitem Input Object

Workitem Input Objects Retrieval for Difference [RO-Q3]

else

Workitem Input Object

Workitem Input Objects Retrieval for Dose [RO-Q4]

Content Assessment Results

Output Info Sequence Attributes

Output Information Sequence Retrieval [RO-Q6]

Quality Check Results Retrieval [RO-Q7]

UPS Create for Quality Check

Quality Check Requester

Quality Check Performer

N-Create (UPS)

Quality Check Requester

Quality Check Performer

Subscribe to UPS Progress Update

Quality Check Requester

Quality Check Performer

N-Action (UPS)

Quality Check Requester

Quality Check Performer

Workitem Input Objects Retrieval

for Difference Check

Quality Check Requester

Quality Check Performer

Retrieve Objects

(

C

-

MOVE

)

Store Objects

(

C

-

STORE

)

Quality Check Performer

Quality Check Requester

Workitem Input Objects Retrieval for Dose Check

Quality Check Requester

Quality Check Performer

Retrieve Objects

(

C

-

MOVE

)

Store Objects

(

C

-

STORE

)

Quality Check Performer

Quality Check Requester

Update on UPS Progress

Quality Check Requester

Quality Check Performer

UPS Event N-EVENT-REPORT

Quality Check Requester

Quality Check Performer

Output Information Sequence Retrieval

Quality Check Requester

Quality Check Performer

N-Get (UPS)

Quality Check Requester

Quality Check Performer

Quality Check Results Retrieval

Quality Check Requester

Quality Check Performer

Quality Check Requester

Quality Check Performer

Store Objects

(

C

-

STORE

)

Move Objects

(

C

-

MOVE

)

Unsubscribe to UPS Progress

Quality Check Requester

Quality Check Performer

N-Action (UPS)

Quality Check Requester

Quality Check Performer